Alyssa Salazar

Bob Alvarenga

Gage Murry

Luis Rodriguez

Mikala Muller

Completed Design Controls Worksheet [DCW-1]

|  |
| --- |
| Comments  Missing: “Improvements on existing design” narrative complete. Design goals clearly defined. (5)   DI-2 easier to assemble is not a binary testable requirement  DI-3 What shall be less than 64 sqcm? The pcb, motor assembly the whole thing including tubing?  DI-5 PCB would be a design output, not a requirement.  25 / 30 pts |
| **Prototyping Protocol [PP-1]** |

|  |
| --- |
| Full Marks  30 / 30 pts |
| **Test Protocol [TP-1]** |

Comments

Compliance Testing: Unclear what the test would be. Does it power up? Function in a specific way? How would the user know what you tested or how you evaluated if it passed or not?  
  
ST1-3 should have bounds on the limits of the values i.e. 0.3 +/- 0.01 mL/s